## Gap assessment ISO13485
CLIENT: |  | <<Client organzations name>> | [8 empty] | DEKRA
ISO 13485:2003 EN ISO 13485:2016 | Gap assessment of (EN) ISO 13485:2016 |  |  |  |  | Implementation |  |  |  |  | Off site assessment |  |  | On site assessment
| CLAUSE assessment |  |  | GAP Assessment |  |  | Evidence that demonstrates implementation of requirements | [6 empty] | effectiveness of implementation
Clause(s) | Clause(s) | Description | New/additional requirements ISO 13485:2016 | GAP | If GAP = Y Indicate your action | Requirement(s) implemented? | Document ID | Document Description | Version | Release date dd-mm-yyyy | Assessed | Acceptable | Remark | Assessed | Acceptable | Evidence assessed / Remark
1 | 1 | Scope | The processes required by this International Standard that are applicable to the organization, but are not performed by the organization, are the responsibility of the organization and are accounted for in the organization’s quality management system by monitoring, maintaining, and controlling the processes.
3 | 3 | Definitions | Several new definitions added and some existing definitions refined.
4 | 4 | Quality management System
4.1 | 4.1 | General requirements
4.1 | 4.1.1 | QMS - general requirements | The organization shall establish, implement and maintain any requirement, procedure, activity or arrangement required to be documented by this International Standard or applicable regulatory requirements.
4.1 | 4.1.1 | QMS - general requirements | The organization shall document the role(s) undertaken by the organization under the applicable regulatory requirements. NOTE Roles undertaken by the organization can include manufacturer, authorized representative, importer or distributor.
4.1 | 4.1.2 | - | b) apply a risk based approach to the control of the appropriate processes needed for the quality management system;
4.1 | 4.1.3 | - | e) establish and maintain records needed to demonstrate conformance to this International Standard and compliance with applicable regulatory requirements (see 4.2.5).
4.1 | 4.1.4 | - | The organization shall manage these quality management system processes in accordance with the requirements of this International Standard and applicable regulatory requirements. Changes to be made to these processes shall be: a) evaluated for their impact on the quality management system; b) evaluated for their impact on the medical devices produced under this quality management system; c) controlled in accordance with the requirements of this International Standard and applicable regulatory requirements.
4.1 | 4.1.5 | - | When the organization chooses to outsource any process that affects product conformity to requirements, it shall monitor and ensure control over such processes. The organization shall retain responsibility of conformity to this International Standard and to customer and applicable regulatory requirements for outsourced processes. The controls shall be proportionate to the risk involved and the ability of the external party to meet the requirements in accordance with 7.4. The controls shall include written quality agreements.
4.1 | 4.1.6 | - | The organization shall document procedures for the validation of the application of computer software used in the quality management system. Such software applications shall be validated prior to initial use and, as appropriate, after changes to such software or its application. The specific approach and activities associated with software validation and revalidation shall be proportionate to the risk associated with the use of the software. Records of such activities shall be maintained (see 4.2.5).
4.2 | 4.2 | Documentation requirements
4.2.1 | 4.2.1 | General | -
4.2.2 | 4.2.2 | Quality manual | -
- | 4.2.3 | Medical device file | For each medical device type or medical device family, the organization shall establish and maintain one or more files either containing or referencing documents generated to demonstrate conformity to the requirement of this International Standard and compliance with applicable regulatory requirements. The content of the file(s) shall include, but is not limited to: a) general description of the medical device, intended use/purpose, and labelling, including any instructions for use; b) specifications for product; c) specifications or procedures for manufacturing, packaging, storage, handling and distribution; d) procedures for measuring and monitoring; e) as appropriate, requirements for installation; f ) as appropriate, procedures for servicing.
4.2.3 | 4.2.4 | Control of documents | g) prevent deterioration or loss of documents;
4.2.4 | 4.2.5 | Control of records | - Addition of: "security and integrity" - The organization shall define and implement methods for protecting confidential health information contained in records in accordance with the applicable regulatory requirements.
4.2.4 | 4.2.5 | Control of records | Changes to a record shall remain identifiable.
5 | 5 | Management responsibility
5.1 | 5.1 | Management commitment
5.1 | 5.1 | Management commitment | -
5.2 | 5.2 | Customer focus
5.1 | 5.1 | Customer focus | -
5.3 | 5.3 | Quality policy
5.1 | 5.1 | Quality policy | -
5.4 | 5.4 | Planning
5.4.1 | 5.4.1 | Quality objectives | -
5.4.2 | 5.4.2 | Quality management system planning | -
5.5 | 5.5 | Responsibility, authority and communication
5.5.1 | 5.5.1 | Responsibility and authority | -
5.5.2 | 5.5.2 | Management representative | c) ensuring the promotion of awareness of applicable regulatory requirements and quality management system requirements throughout the organization.
5.5.3 | 5.5.3 | Internal communication | -
5.6 | 5.6 | Management review
5.6.1 | 5.6.1 | General | The organization shall document procedures for management review.
5.6.2 | 5.6.2 | Review input | b) complaint handling c) reporting to regulatory authorities; l) applicable new or revised regulatory requirements.
5.6.3 | 5.6.3 | Review output | The output shall be recorded and include input reviewed c) changes needed to respond to applicable new or revised regulatory requirements;
6 | 6 | Resource management
6.1 | 6.1 | Provision of resources
6.1 | 6.1 | Provision of resources | -
6.2 | 6.2 | Human resources
6.2 | 6.2 | Human resources | The organization shall document the process(es) for establishing competence, providing needed training, and ensuring awareness of personnel. The methodology used to check effectiveness is proportionate to the risk associated with the work for which the training or other action is being provided.
6.3 | 6.3 | Infrastructure
6.3 | 6.3 | Infrastructure | The organization shall document the requirements for the infrastructure needed to achieve conformity to product requirements, prevent product mix-up and ensure orderly handling of product. The organization shall document requirements for the maintenance activities, including the interval of performing the maintenance activities, when such maintenance activities, or lack thereof, can affect product quality. As appropriate, the requirements shall apply to equipment used in production, the control of the work environment and monitoring and measurement.
6.4 | 6.4 | Work environment and contamination control
6.4 | 6.4.1 | Work environment | -
6.4 | 6.4.2 | Contamination control | For sterile medical devices, the organization shall document requirements for control of contamination with microorganisms or particulate matter and maintain the required cleanliness during assembly or packaging processes.
7 | 7 | Product realization
7.1 | 7.1 | Planning of product realization
7.1 | 7.1 | Planning of product realization | The organization shall document one or more processes for risk management in product realization. Records of risk management activities shall be maintained (see 4.2.5). b) the need to establish processes and documents (see 42A) and to provide resources specific to the product, including infrastructure and work environment; c) required verification, validation, monitoring, measurement, inspection and test, handling, storage, distribution and traceability activities specific to the product together with the criteria for product acceptance;
7.2 | 7.2 | Customer-related processes
7.2.1 | 7.2.1 | Determination of requirements related to product | d) any user training needed to ensure specified performance and safe use of the medical device;
7.2.2 | 7.2.2 | Review of requirements related to product | c) applicable regulatory requirements are met; d) any user training identified in accordance with 7.2.1 is available or planned to be available;
7.2.3 | 7.2.3 | Communication | The organization shall communicate with regulatory authorities in accordance with applicable regulatory requirements.
7.3 | 7.3 | Design and development
7.3.1 | 7.3.1 | General | -
7.3.1 | 7.3.2 | Design and development planning | During design and development planning, the organization shall document: e) the methods to ensure traceability of design and development outputs to design and development inputs; f) the resources needed, including necessary competence of personnel.
7.3.2 | 7.3.3 | Design and Development inputs | a) functional, performance, usability and safety requirements, according to the intended use; - Requirements shall be complete, unambiguous, able to be verified or validated, and not in conflict with each other.
7.3.3 | 7.3.4 | Design and development outputs | -
7.3.4 | 7.3.5 | Design and Development review | At suitable stages, systematic reviews of design and development shall be performed in accordance with planned and documented arrangements to: Records of the results of the reviews and any necessary actions shall be maintained and include the identification of the design under review, the participants involved and the date of the review (see 4.2.5).
7.3.5 | 7.3.6 | Design and Development verification | The organization shall document verification plans that include methods, acceptance criteria and, as appropriate, statistical techniques with rationale for sample size. If the intended use requires that the medical device be connected to, or have an interface with, other medical device(s), verification shall include confirmation that the design outputs meet design inputs when so connected or interfaced. Records of the results and conclusions of the verification and necessary actions shall be maintained (see 42A and 4ZS).
7.3.6 | 7.3.7 | Design and Development validation | - The organization shall document validation plans that include methods, acceptance criteria and, as appropriate, statistical techniques with rationale for sample size. - Design validation shall be conducted on representative product. Representative product includes initial production units, batches or their equivalents. The rationale for the choice of product used for validation shall be recorded (see 4.2.5). - As part of design and development validation, the organization shall perform clinical evaluations or performance evaluations of the medical device in accordance with applicable regulatory requirements. A medical device used for clinical evaluation or performance evaluation is not considered to be released for use to the customer. - If the intended use requires that the medical device be connected to, or have an interface with, other medical device(s), validation shall include confirmation that the requirements for the specified application or intended use have been met when so connected or interfaced. Records of the results and conclusion of validation and necessary actions shall be maintained (see 42A and4).
- | 7.3.8 | Design and Development transfer | The organization shall document procedures for transfer of design and development outputs to manufacturing. These procedures shall ensure that design and development outputs are verified as suitable for manufacturing before becoming final production specifications and that production capability can meet product requirements. Results and conclusions of the transfer shall be recorded (see 4.2.5).
7.3.7 | 7.3.9 | Control of design and development changes | The organization shall document procedures to control design and development changes. The organization shall determine the significance of the change to function, performance, usability, safety and applicable regulatory requirements for the medical device and its intended use. - The review of design and development changes shall include evaluation of the effect of the changes on constituent parts and product in process or already delivered, inputs or outputs of risk management and product realization processes.
- | 7.3.10 | Design and Development files | The organization shall maintain a design and development file for each medical device type or medical device family. This file shall include or reference records generated to demonstrate conformity to the requirements for design and development and records for design and development changes.
| 7.4.1 | Purchasing process | The organization shall establish criteria for the evaluation and selection of suppliers. The criteria shall be: a) based on the supplier’s ability to provide product that meets the organization’s requirements; b) based on the performance of the supplier; c) based on the effect of the purchased product on the quality of the medical device; d) proportionate to the risk associated with the medical device. The organization shall plan the monitoring and re-evaluation of suppliers. Supplier performance in meeting requirements for the purchased product shall be monitored. The results of the monitoring shall provide an input into the supplier re-evaluation process. Non-fulfilment of purchasing requirements shall be addressed with the supplier proportionate to the risk associated with the purchased product and compliance with applicable regulatory requirements. Records of the results of evaluation, selection, monitoring and re-evaluation of supplier capability orperformance and any necessary actions arising from these activities shall be maintained (see 4.2.5).
7.4 | 7.4 | Purchasing
7.4.1 | 7.4.1 | Purchasing process | The organization shall establish criteria for the evaluation and selection of suppliers. The criteria shall be: a) based on the supplier’s ability to provide product that meets the organization’s requirements; b) based on the performance of the supplier; c) based on the effect of the purchased product on the quality of the medical device; d) proportionate to the risk associated with the medical device. The organization shall plan the monitoring and re-evaluation of suppliers. Supplier performance in meeting requirements for the purchased product shall be monitored. The results of the monitoring shall provide an input into the supplier re-evaluation process. Non-fulfilment of purchasing requirements shall be addressed with the supplier proportionate to the risk associated with the purchased product and compliance with applicable regulatory requirements.
7.4.2 | 7.4.2 | Purchasing information | a) product specifications; b) requirements for product acceptance, procedures, processes and equipment; - Purchasing information shall include, as applicable, a written agreement that the supplier notify the organization of changes in the purchased product prior to implementation of any changes that affect the ability of the purchased product to meet specified purchase requirements.
7.4.3 | 7.4.3 | Verification of purchased product | The extent of verification activities shall be based on the supplier evaluation results and proportionate to the risks associated with the purchased product. When the organization becomes aware of any changes to the purchased product, the organization shall determine whether these changes affect the product realization process or the medical device.
7.5 | 7.5 | Production and service provision
7.5.1 7.5.1.1 | 7.5.1 | Control of production and service provision | b) qualification of infrastructure; c) implementation of monitoring and measurement of process parameters and product characteristics;
7.5.1.2.1 | 7.5.2 | Cleanliness of product | c) product cannot be cleaned prior to sterilization or its use, and its cleanliness is of significance in use;
7.5.1.2.2 | 7.5.3 | Installation activities | -
7.5.1.2.3 | 7.5.4 | Servicing activities | The organization shall analyse records of servicing activities carried out by the organization or its supplier: a) to determine if the information is to be handled as a complaint; b) as appropriate, for input to the improvement process.
7.5.1.3 | 7.5.5 | Particular requirements for sterile medical devices | -
7.5.2 | 7.5.6 | Validation of processes for production and service provision | The organization shall validate any processes for production and service provision where the resultingoutput cannot be or is not verified by subsequent monitoring or measurement and, as a consequence, deficiencies become apparent only after the product is in use or the service has been delivered. Validation shall demonstrate the ability of these processes to achieve planned results consistently. The organization shall document procedures for validation of processes, including: d) as appropriate, statistical techniques with rationale for sample sizes; g) approval of changes to the processes. The organization shall document procedures for the validation of the application of computer software used in production and service provision. Such software applications shall be validated prior to initial use and, as appropriate, after changes to such software or its application. The specific approach and activities associated with software validation and revalidation shall be proportionate to the risk associated with the use of the software, including the effect on the ability of the product to conform to specifications. Records of the results and conclusion of validation and necessary actions from the validation shall be maintained (see 4.2.4 and 4.2.5).
7.5.2.2 | 7.5.7 | Particular requirements for validation of processes for sterilization and sterile barrier systems | The organization shall document procedures (see 42A) for the validation of processes for sterilization and sterile barrier systems. Processes for sterilization and sterile barrier systems shall be validated prior to implementation and following product or process changes, as appropriate. Records of the results and, conclusion of validation and necessary actions from the validation shall be maintained (see 42A and 4.2.5).
7.5.3 7.5.3.3 | 7.5.8 | Identification | The organization shall identify product status with respect to monitoring and measurement requirements throughout product realization. Identification of product status shall be maintained throughout production, storage, installation and servicing of product to ensure that only product that has passed the required inspections and tests or released under an authorized concession is dispatched, used or installed. If required by applicable regulatory requirements, the organization shall document a system to assign unique device identification to the medical device.
7.5.3.2 | 7.5.9 | Traceability | -
7.5.3.2.1 | 7.5.9.1 | General | The organization shall document procedures for traceability. These procedures shall define the extent of traceability in accordance with applicable regulatory requirements and the records to be maintained (see 4.2.5).
7.5.3.2.2 | 7.5.9.2 | Particular requirements for implantable medical devices | -
7.5.4 | 7.5.10 | Customer property | -
7.5.5 | 7.5.11 | Preservation of product | The organization shall protect product from alteration, contamination or damage when exposed to expected conditions and hazards during processing, storage, handling, and distribution by: a) designing and constructing suitable packaging and shipping containers; b) documenting requirements for special conditions needed if packaging alone cannot provide preservation. If special conditions are required, they shall be controlled and recorded (see 4.2.5).
7.6 | 7.6 | Control of monitoring and measuring equipment | b) be adjusted or re-adjusted as necessary: such adjustments or re-adjustments shall be recorded (see 4.2.5); c) have identification in order to determine its calibration status; The organization shall document procedures for the validation of the application of computer software used for the monitoring and measurement of requirements. Such software applications shall be validated prior to initial use and, as appropriate, after changes to such software or its application. The specific approach and activities associated with software validation and revalidation shall be proportionate to the risk associated with the use of the software, including the effect on the ability of the product to conform to specifications. Records of the results and conclusion of validation and necessary actions from the validation shall be maintained (see 4.2.4 and 4.2.5).
8 | 8 | Measurement, analysis and improvement
8.1 | 8.1 | General
8.1 | 8.1 | General | -
8.2 | 8.2 | Monitoring and measurement
8.2.1 | 8.2.1 | Feedback | The methods for obtaining and using this information shall be documented. The organization shall document procedures for the feedback process. This feedback process shall include provisions to gather data from production as well as post-production activities. The information gathered in the feedback process shall serve as potential input into risk management for monitoring and maintaining the product requirements as well as the product realization or improvement processes. If applicable regulatory requirements require the organization to gain specific experience from postproduction activities, the review of this experience shall form part of the feedback process.
7.2.3 8.2.1 | 8.2.2 | Complaint handling | The organization shall document procedures for timely complaint handling in accordance with applicable regulatory requirements. These procedures shall include at a minimum requirements and responsibilities for: a) receiving and recording information; b) evaluating information to determine if the feedback constitutes a complaint; c) investigating complaints; d) determining the need to report the information to the appropriate regulatory authorities; e) handling of complaint-related product; f) determining the need to initiate corrections or corrective actions. If any complaint is not investigated, justification shall be documented. Any correction or corrective action resulting from the complaint handling process shall be documented. If an investigation determines activities outside the organization contributed to the complaint, relevant information shall be exchanged between the organization and the external party involved. Complaint handling records shall be maintained (see 4.2.5).
- | 8.2.3 | Reporting to regulatory authorities | If applicable regulatory requirements require notification of complaints that meet specified reporting criteria of adverse events or issuance of advisory notices, the organization shall document procedures for providing notification to the appropriate regulatory authorities. Records of reporting to regulatory authorities shall be maintained (see 4.2.5).
8.2.2 | 8.2.4 | Internal audit | a) conforms to planned and documented arrangements, requirements of this International Standard, quality management system requirements established by the organization, and applicable regulatory requirements; The audit criteria, scope, interval and methods shall be defined and recorded (see 4.2.5).
8.2.3 | 8.2.5 | Monitoring and measurement of processes | The organization shall apply suitable methods for monitoring and, as appropriate, measurement of the quality management system processes. These methods shall demonstrate the ability of the processes to achieve planned i’esults. When planned results are not achieved, correction and corrective action shall be taken, as appropriate. - REMOVED to ensure conformity of the product REMOVED -
8.2.4 8.2.4.1 8.2.4.2 | 8.2.6 | Monitoring and measurement of product | As appropriate, records shall identify the test equipment used to perform measurement activities.
8.3 | 8.3 | Control of nonconforming product
8.3 | 8.3.1 | General | The organization shall document a procedure to define the controls and related responsibilities and authorities for the identification, documentation, segregation, evaluation and disposition of nonconforming product. The evaluation of nonconformity shall include a determination of the need for an investigation and notification of any external party responsible for the nonconformity. Records of the nature of the nonconformities and any subsequent action taken, including the evaluation, any investigation and the rationale for decisions shall be maintained (see 4.2.5)
8.3 | 8.3.2 | Actions in response to nonconforming product detected before delivery | The organization shall ensure that nonconforming product is accepted by concession only if the justification is provided, approval is obtained and applicable regulatory requirements are met. Records of the acceptance by concession and the identity of the person authorizing the concession shall bemaintained (see 4.2.5).
8.3 | 8.3.3 | Actions in response to nonconforming product detected after delivery | Records of actions taken shall be maintained (see 4.2.5). The organization shall document procedures for issuing advisory notices in accordance with applicable regulatory requirements. These procedures shall be capable of being put into effect at any time. Records of actions relating to the issuance of advisory notices shall be maintained (see 4.2.5).
8.3 | 8.3.4 | Rework | After the completion of rework, product shall be verified to ensure that it meets applicable acceptance criteria and regulatory requirements. Records of rework shall be maintained (see 4.2.5).
8.4 | 8.4 | Analysis of data
8.4 | 8.4 | Analysis of data | The procedures shall include determination of appropriate methods, including statistical techniques and the extent of their use. e) audits; f) service reports, as appropriate. If the analysis of data shows that the quality management system is not suitable, adequate or effective, the organization shall use this analysis as input for improvement as required in 8.5.
8.5 | 8.5 | Improvement
8.5.1 | 8.5.1 | General | The organization shall identify and implement any changes necessary to ensure and maintain the continued suitability, adequacy and effectiveness of the quality management system as well as medical device safety and performance through the use of the quality policy, quality objectives, audit results, postmarket surveillance, analysis of data, corrective actions, preventive actions and management review.
8.5.2 | 8.5.2 | Corrective actions | (..) Any necessary corrective actions shall be taken without undue delay. Corrective actions shall be proportionate to the effects of the nonconformities encountered. The organization shall document a procedure to define requirements for: d) planning and documenting action needed and implementing such action, including, as appropriate, updating documentation; e) verifying that the corrective action does not adversely affect the ability to meet applicable regulatory requirements or the safety and performance of the medical device;
8.5.3 | 8.5.3 | Preventive actions | Preventive actions shall be proportionate to the effects of the potential problems. The organization shall document a procedure to describe requirements for: c) planning and documenting action needed and implementing such action, including, as appropriate, updating documentation; d) verifying that the action does not adversely affect the ability to meet applicable regulatory requirements or the safety and performance of the medical device;
END
1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1

## Combination with ISO 9001
CLIENT: | <<Client organzations name>> | [8 empty] | DEKRA
Assessment combination with ISO 9001:2015 |  |  |  |  | Implementation |  |  |  |  | Off site assessment |  |  | On site assessment
CLAUSE assessment |  |  | GAP Assessment |  |  | Evidence that demonstrates implementation of requirements | [6 empty] | effectiveness of implementation
Clause(s) | Description | Additional requirements ISO 9001:2015 | GAP | If GAP = Y Indicate your action | Requirement(s) implemented? | Document ID | Document Description | Version | Release date dd-mm-yyyy | Assessed | Acceptable | Remark | Assessed | Acceptable | Evidence assessed / Remark
4. | Context of the organization
4.1 | Understanding the organization and its context | The organization shall determine external and internal issues that are relevant to its purpose and its strategic direction and that affect its ability to achieve the intended result(s) of its quality management system. The organization shall monitor and review information about these external and internal issues.
4.2 | Understanding the needs and expectations of interested parties | Due to their effect or potential effect on the organization’s ability to consistently provide products and services that meet customer and applicable statutory and regulatory requirements, the organization shall: Determine the interested parties that are relevant to the quality management system. Determine the requirements of these interested parties that are relevant to the quality management system. The organization shall monitor and review information about these interested parties and their relevant requirements.
6 | Planning
6.1 | Actions to address risks and opportunities
6.1.1 | Risks and opportunities | When planning for the quality management system, the organization shall consider the issues referred to in 4.1 and the requirements referred to in 4.2 and determine the risks and opportunities that need to be addressed to: a) give assurance that the quality management system can achieve its intended result(s); b) enhance desirable effects; c) prevent, or reduce, undesired effects; d) achieve improvement.
6.1.2 | Risks and opportunities | The organization shall plan: a) actions to address these risks and opportunities; b) how to: 1) integrate and implement the actions into its quality management system processes (see 4.4); 2) evaluate the effectiveness of these actions. Actions taken to address risks and opportunities shall be proportionate to the potential impact on the conformity of products and services.
6.2 | Quality objectives and planning to achieve them
6.2.2 | Quality objectives | When planning how to achieve its quality objectives, the organization shall determine: a) what will be done; b) what resources will be required; c) who will be responsible; d) when it will be completed; e) how the results will be evaluated.
7 | Support
7.4 | Communication
7.4 | Communication | The organization shall determine the internal and external communications relevant to the quality management system, including: a) on what it will communicate; b) when to communicate; c) with whom to communicate; d) how to communicate; e) who communicates.
5.4.2 | Quality management system planning | -
9 | Performance evaluation
9.1 | Monitoring, measurement, analysis and evaluation
9.1.2 | Customer satisfaction | The organization shall monitor customers’ perceptions of the degree to which their needs and expectations have been fulfilled. The organization shall determine the methods for obtaining, monitoring and reviewing this information.
9.3 | Management review
9.3.2 | Management review inputs | The management review shall be planned and carried out taking into consideration: d) the adequacy of resources; e) the effectiveness of actions taken to address risks and opportunities (see 6.1);
END
1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1

## Subclause matrix
Client Job - report # | ISO 13485:2016 | 4.1 General requirements | 4.1.4 Changes | 4.1.5 Outsourced processes | 4.1.6 - Software validation QMS | 4.2 Documentation requirements 4.2.1 General. | 4.2.2 Quality manual | 4.2.3 Medical device file | 4.2.4 Control of documents | 4.2.5 Control of records | 5.1 – management commitment | 5.2 – Customer focus | 5.3 – Quality policy | 5.4.1 – Quality objectives | 5.4.2 – QMS planning | 5.5.1 – Responsibility and authority | 5.5.2 – Management representative | 5.5.3 – Internal communication | 5.6.1 – Management review | 5.6.2 – Review input | 5.6.3 – Review output | 6.1 – Provision of resources | 6.2.1 – Human resources | 6.2.2 - Competence, awareness and training | 6.3 – Infrastructure | 6.4 - Work environment | 6.4.1 Work environment | 6.4.2 Contamination control | 7.1 – Planning of product realization | 7.2.1 Determination of requirements related to product | 7.2.2 Review of requirements related to product | 7.2.3 Communication | 7.3.1 - Design and development | 7.3.2 Design and development planning | 7.3.3 Design and development inputs | 7.3.4 Design and development outputs | 7.3.5 Design and development review | 7.3.6 Design and development verification | 7.3.7 Design and development validation | 7.3.8 Design and development transfer | 7.3.9 Control of design and development changes | 7.3.10 Design and development files | 7.4.1 – Purchasing process | 7.4.2 Purchasing information | 7.4.3 Verification of purchased product | 7.5.1 Control of production and service provision | 7.5.2 Cleanliness of product | 7.5.3 Installation activities | 7.5.4 Servicing activities | 7.5.5 Particular requirements for sterile medical devices | 7.5.6 Validation of processes for production and service provision | 7.5.7 Particular requirements for validation of processes for sterilization and sterile barrier systems | 7.5.8 Identification | 7.5.9.1 Traceability | 7.5.9.2 Particular requirements for implantable medical devices | 7.5.10 Customer property | 7.5.11 - Preservation of product | 7.6 Control of monitoring and measuring equipment | 8.1 Measurement, analysis and improvement | 8.2.1 Feedback | 8.2.2 Complaint handling | 8.2.3 Reporting to regulatory authorities | 8.2.4 - Internal auditing | 8.2.5 Monitoring and measurement of processes | 8.2.6 - Monitoring and measurement of product | 8.3 - Control of nonconforming product | 8.4 - Analysis of data | 8.5.1 Improvement | 8.5.2 Corrective action | 8.5.3 Preventive action | Check
Remove row in case ISO 9001 is not applicable | Combination with ISO 9001:2015 | 4.1 Understanding the organization and its context 4.2 Understanding the need and expectations of interested parties | 6.3 Planning of changes 8.5.6 Control of changes | 8.4.3 Information for external providers | 8.5.1 Control of production and service provision | 7.5.1 General | 4.3 Determining the scope of the QMS 4.4.1 QMS and its processes | 4.4.2 Documented information | 7.5.3 Control of documented information | 7.5.3 Control of documented information | 5.1 Leadership and commitment | 5.1.2 Customer focus | 5.2.1 Establishing the quality policy 5.2.2 Communicating the quality policy | 6.2 Quality objectives and planning to achieve them | 6.3 Planning of changes 6.1 Actions to adress risks and opportunities | 4.4.1.e responsibilities and authorities | 5.3 Organizational roles, responsibilities and authorities | 7.4 Communication | 9.3.1 General (9.3 Management review) | 9.3.2 Management review inputs | 9.3.3 Management review outputs | 7.1.2 People | 7.1.6 Organizational knowledge 7.2 Competence 7.3 Awareness | 7.1.6 Organizational knowledge 7.2 Competence 7.3 Awareness | 7.1.3 Infrastructure | 7.1.4 Environment for the operation of processes | 7.1.4 Environment for the operation of processes | 7.1.4 Environment for the operation of processes | 8.1 Operational planning and control | 8.2.2 Determining the requirements for products and services | 8.2.3.1 Review of the requirements for products and services | 7.4 Communication 8.2.1 Customer communication | 8.3.1 Design and development | 8.3.2 Design and development planning | 8.3.3 Design and development inputs | 8.3.5 Design and development outputs | 8.3.4 Design and development controls | 8.3.4 Design and development controls | 8.3.4 Design and development controls | 8.3.5 Design and development outputs | 8.3.6 Design and development changes | 8.3.5 Design and development outputs | 8.4.1 Control of externally provided processes, products and services | 8.4.3 Information for external providers | 8.4.2 Type and extent of control | 8.5.1 Control of production and service provision | 8.5.1 Control of production and service provision | 8.5.5 Post-delivery activities | 8.5.1 Control of production and service provision 8.5.5 Post-delivery activities | 8.5.1 Control of production and service provision | 7.1.5.1 General - 7.1.5 Monitoring and measurment resources 8.5.1 Control of production and service provision | 8.5.1 Control of production and service provision | 8.5.2 Identification and traceability | 8.5.2 Identification and traceability | 8.5.2 Identification and traceability | 8.5.3 Property belonging to customers or external providers | 8.5.4 Preservation | 7.1.5 Monitoring and measuring resources | 9.1 Monitoring, measurment, analysis and evaluation | 9.1.2 Customer satisfaction | 9.1.2 Customer satisfaction | 8.5.5 Post-delivery activities 9.1.2 Customer satisfaction | 9.2 Internal audit | 9.1 Monitoring, measurment, analysis and evaluation | 8.6 Release of products and services | 8.7 Control of nonconforming outputs | 9.1.3 Analysis and evaluation | 6.3 Planning of changes 8.5.6 Control of changes | 10.2 Nonconformity and corrective action | 10.1 General - 10 Improvement 6.1 Actions to address risks and opportunities
Continual Management Dynamics (e.g., 1.03, 1.08, 1.02) | [70 empty] | 0
On site audit (Upgrade) | [70 empty] | 0
Management Processes | [70 empty] | 0
Management System Document Review | [70 empty] | 0
Management Involvement and Customer Focus includes: | [70 empty] | 0
- Policy and Objectives | [70 empty] | 0
- Quality Planning | [70 empty] | 0
- Responsibility, Authority and Communication | [70 empty] | 0
Management Review
Document and Data Control | [70 empty] | 0
Resource Management (people, machines, work areas) | [70 empty] | 0
Competence and Qualification | [70 empty] | 0
Facilities Control | [70 empty] | 0
Calibration and Preventive Maintenance | [70 empty] | 0
Computer Systems Management | [70 empty] | 0
Customer Processes | [70 empty] | 0
Contract Review (Orders Processing) | [70 empty] | 0
Returns Goods Handling and Processing | [70 empty] | 0
Distributor Control | [70 empty] | 0
Post Market Surveillance | [70 empty] | 0
Product Realization Processes (should include training, equipment, data analysis and records control associated with these processes) | [70 empty] | 0
Purchasing & Supplier Management | [70 empty] | 0
Receiving/Receiving Inspection | [70 empty] | 0
Production Planning | [70 empty] | 0
Product Realization – process A | [70 empty] | 0
Product Realization – process B | [70 empty] | 0
Product Realization – Testing | [70 empty] | 0
Final Inspection and Release | [70 empty] | 0
Warehousing/ Inventory Control/ Shipping | [70 empty] | 0
Improvement Processes | [70 empty] | 0
Nonconforming Product | [70 empty] | 0
Corrective and Preventive Action | [70 empty] | 0
Internal Audits | [70 empty] | 0
Complaints, Vigilance, Recall | [70 empty] | 0
Post Market Surveillance | [70 empty] | 0
0
Technical Documentation Review | [70 empty] | 0
Clinical evaluation | [70 empty] | 0
Annex II capability | [70 empty] | 0
Certification Update | [70 empty] | 0
0
Medical device licensing for Canada | [70 empty] | 0
0
Taiwanese TCP overview | [70 empty] | 0
0
…………………… | [70 empty] | 0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
|  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0

## Version
Version | Change | Date
0 | Initial version | 2017-03-17
1 | Work sheets updated to enable inserting rows and formatting cells that are not protected. Can be used in case multiple documents require to be entered to demonstrate compliance to the applicable requirement(s). | 2017-04-04
2 | 1. Added On site assessment - effectiveness of implementation in columns N; O and P to enable full assessment within this document for both sheets Gap assessment ISO13485 and Combination with ISO9001. 2. Added Subclause matrix as separate sheet with row 4 "On site audit (Upgrade)" auto populated. | 2017-07-19
3 | Updated sheet Subclause matrix to enable adding and deleting rows and columns | 2017-08-15